## Slide 1: SPE153 CBD
Feasibility study on automation of monthly reporting
28th Sept 2023
Lloyd Clayburn - Healthcare Scientist Trainee (STP) lloyd.clayburn@nhs.net
Notes: 15 m
--
MG suggests that GPT4 API (~£10/mth) could be relatively easily programmed to do a better job
PV says Computer Vision Toolbox is ~£100, TV might be put to work on this.

## Slide 2
Introduction
Optical character recognition
Monthly reporting
Literature
Can we automate monthly reporting?
Code MATLAB demonstration
Development
Results and discussion
Further work
Summary

## Slide 3
Monthly reporting
Reverberation depth, gain, colour gain
Not included in DICOM
Advisory service for 79 scanners across NGH, RHH, JWH
Something
Notes: Advisory service
Est ~ 3 weeks/year reporting monthly QA
--
Scanner count:
Ngh 15 7 5 1 1 10
Rhh 2 2 4 4 2 2 7
jwh 4 2 8 3
39 23 17 = 79
Time est:
If 1 h for scan dept (8)
10 h/mth for all
120 h / year = 3 b6/b7 weeks/year

## Slide 4
Optical character recognition
Machine learning application
MATLAB ocr() uses open-source C++ Tesseract code [1]
Introduced MATLAB 2014a: part of computer vision toolbox [1]
Mathworks.com.
Notes: Developed by HP 1985-1995, refined and released open source in 2005
2006-2018 developed by Google
Available in python as pytesseract()
--
MATLAB – introduced in 2014a [1]
Part of computer vision toolbox [1]
Uses open source C++ Tesseract code [2]
Here it is https://github.com/tesseract-ocr/ [3]
Developed by HP 1985-1995, refined and released open source in 2005 [4]
2006-2018 developed by Google [3]
Initially trained on 20 samples of 94 characters from 8 fonts in a single size, but with 4 attributes (normal, bold, italic, bold italic) [4]
Available in python as pytesseract() [5]
1. OCR [Internet] Mathworks 2023 [cited 2023 Sept 28]. Available from: https://uk.mathworks.com/help/vision/ref/ocr.html
2. Getting started with OCR [Internet] Mathworks 2023 [cited 2023 Sept 28]. Available from: https://uk.mathworks.com/help/vision/ug/getting-started-with-ocr.html
3. Tesseract – OCR [Internet] Available from: https://github.com/tesseract-ocr/
4. R. Smith. An Overview of the Tesseract OCR Engine. Ninth International Conference on Document Analysis and Recognition (ICDAR), Curitiba, Brazil, 2007, pp. 629-633, doi: 10.1109/ICDAR.2007.4376991.
5. pytesseract 0.3.10 [Internet]. 2022 [cited 2023 Sept 28]. Available from: https://pypi.org/project/pytesseract/

## Slide 5
US image OCR & QC automation
Reul et al. 2016 Evaluation of OCR on US DICOM.
Horssen et al. 2017 Automated Quantitative QC for in-air reverb images.
Stekel et al. 2019 OCR on tens of thousands of US DICOM for analytics.
Notes: Reul et al 2016
OCR on US DICOM. Open-source java code inc. TesseractOCR.
Interesting stuff that my code is missing:
ROI identification (text is usually on background with grayscale value x)
Segmentation of text containing regions into lines.
Feed OCR line by line.
Horssen 2017
Quantitative QC based on in-air reverb images.
These images serve as an initial indication of image degradation.
Automated analysis extracts quality parameters.
Static images acquired by the clinical user are transferred to a server where analysis is performed.
Stekel et al 2019
OCR on tens of thousands of US DICOM for analytics
Gives example of auditing preset checking with info
How much are presets used once introduced?
Trained on specific fonts used by scanner manufacturers
LC – TI and MI could be audited in this way
--
Reul C, Köberle P, Üçeyler N, Puppe F. Expectation-Driven Text Extraction from Medical Ultrasound Images. Stud Health Technol Inform. 2016;228:712-6.
van Horssen P, Schilham A, Dickerscheid D et al. Automated quality control of ultrasound based on in-air reverberation patterns. Ultrasound. 2017 Nov;25(4):229-238. doi: 10.1177/1742271X17733145. Epub 2017 Sep 29.
Stekel SF, Long Z, Tradup DJ, Hangiandreou NJ. Use of Image-Based Analytics for Ultrasound Practice Management and Efficiency Improvement. J Digit Imaging. 2019 Apr;32(2):251-259. doi: 10.1007/s10278-018-0151-2.

## Slide 6
Can we automate monthly reporting?
Download compressed DICOM from PACS
Place in relevant folders on QAImageStore drive
Tidy up as downloaded files – LC MATLAB script
Decompress DICOM – Scientific Computing MATLAB script
Copy QA data from images into .xls – OCR MATLAB script
Check values in tolerance, bold and comment if not
Check images for changes between this QA & last, comment
Fill out report
Notes: Separate fn for file restructuring - not discussed here

## Slide 7
Code
Initialise parameters
Get scanner ID from directory, model & study date from header
Check for reseats and remove from consideration
Error if number of files differs from that expected
For each probe:
Get Probe Number
Get QACLIN reverb depth, gain colour gain
if error, try with different parameters
if still error, prompt for manual input
Get QAPHYS reverb depth, gain, colour gain
Append results to .xls
https://github.com/ClayburnL/USQA
Notes: Initialise parameters: invert, binarise, layout
Current system relies on reliable file structure - error if more or less files than expected

## Slide 8: MATLAB demonstration

## Slide 9
Development
Initial testing & development on Mar 2023 JWH Scan Dept data
3 GE Voluson E8 1 Philips IU22 4 Toshiba Aplio MX
Further testing & development on Feb data
Notes: March data
Earlier iteration achieved 100% accuracy on March data
Feb data
Missing zero, missing d.p.s, incorrect readings among errors
Had to clean duplicate images for testing
Cine review pane caused bother

## Slide 10
Results – test on Jan 2023 data
Scanner | Errors caught | Errors missed | Accuracy | Comments
GE_2018-00856 | 0 | 0 | 21/21 | 
GE_2018-00857 | 0 | 0 | 21/21 | 
GE_2018-00858 | 0 | 0 | 21/21 | Failed initially - fixed bug
IU22_2009-10800 |  |  |  | Incorrect number of files
Toshiba_2011-02498 |  |  |  | Incorrect number of files
Toshiba_2011-02499 | 0 | 1 | 20/21 | Incorrect colour gain
Toshiba_2012-00122 | 0 | 2 | 19/21 | 2x incorrect probe ID
Toshiba_2012-00123 | 0 | 6 | 15/21 | Missing figures after d.p.
Notes: Counting manual data inputs as successes
Evaluation on v17 200923
OCR accuracy evaluated through comparison between manual observer and .xls
21 numbers to read per scanner
Caught error would be flagged as such
IU22 error due to incorrect number of files resolved following testing by moving append2csv into ‘if’ statement

## Slide 11
Further work
Code improvements
Extend to more scanners?
Scientific computing interested to progress?
Flag utility of DICOM tags for QC params to manufacturers?
Notes: Code improvements
Add confidence level checking and missing data catching to ‘get probe number’
Get iteration through file structure working
Standardise error catching and attempts at resolution through code modules
Catch duplicate images? - e.g. by recognising reverb has dist text, CG has CG data... - big job
Extend to more scanners?
Automate ROI ID and text segmentation?
Scientific computing interested to progress?
  Improve robustness, pythonise
  Automate other parts of the chain: Include DICOM decompress in file structuring code
  Read csv and flag up out of tolerance values

## Slide 12
Summary
Tesseract OCR engine implemented as MATLAB ocr()
Used to read monthly QA values from DICOM for JWH Scan Dept.
118/126 (94%) accuracy in where no error due to additional images

## Slide 13: Thanks for listening
Lloyd Clayburn - Healthcare Scientist Trainee (STP) lloyd.clayburn@nhs.net